## URD + MCRD Requirements
URD # | Req # | Requirement Description | Priority (1,2,3) | PRD #
| URD-U1 | The User should be able to use this device as their primary speech generating device in their home and while in transit, in a maximum variety of contexts and lighting conditions
URD-U01 |  | The User should be able to use this device in a variety of contexts and lighting conditions
| 8.1.6 | The screen on the system should be legible to the user in an indoor and outdoor setting. | 2
| 8.6.1 | The system should generate speech even when not connected to the internet. | 1
| 8.6.2 | The system should be able to access banked voices even when not connected to the internet. | 2
| 8.6.7 | They system should not require cellular data to operate enable use of a cell-phone as a hot-spot for internet access. | 1
URD-U02 |  | The User should be able to use this device for at least 60 consecutive minutes to generate speech
| 8.1.1 | The system should be comfortable to wear for 1 – 2 hours. | 1
| 8.1.2 | The system should be comfortable to wear for 6-8 hours. | 3
| 8.1.4 | The system should have a self-contained battery life of 2 hours | 1
| 8.1.5 | The system should provide an extended battery option for 8 hours of standard use. | 3
| 8.1.11 | The use of the headset should be a pleasant and relaxing experience both from a feel of the headset and the design of the user interface | 2
| 8.10.1 | The system should provide ample warning of low and critical battery status. indicate low battery status when 30 minutes of standard use remain and critical battery status when 10 minues remain. | 2
| URD-U2 | The headset should accomodate a maximum types of user bodies and environments
URD-U03 |  | The headset should accommodate more than 75% of user heads
| 8.1.3 | The system should fit at least 95% of average adult heads (i.e. within normal limits). | 2
URD-U04 |  | The headset should accommodate use in a variety of user positions (such as in a wheelchair or in a bed)
| 8.1.9 | The system should be comfortable in a seated and reclined position. (e.g. when a patient with ALS wakes up, they should be able to communicate an awakened state to their caregiver) | 2
| 8.1.10 | The system should be comfortable for use in a wheelchair with a headrest. | 1
URD-U05 |  | The headset should accommodate users who need hearing aids or glasses
| 8.2.10 | The system should be adaptable for users who wear glasses | 2
| 8.2.11 | The system should be adaptable for users with hearing peripherals | 3
URD-U06 | URD-U3 | The headset should accommodate a variety of input modes based on user’s physical ability
| 8.2.1 | The system should allow for input modes of head movement and visual evoked potentials. | 1
| 8.2.2 | The system should allow for input modes of eye tracking | 2
| 8.2.3 | The system should allow the user physical methods such as a joystick or mouse to focus on a visual target (instead of head pose, for example) | 3
| 8.2.4 | The system should enable the use of a switch for quicker selection regardless of input mode being used. | 2
| 8.2.5 | The system should be able to measure user’s ability to use the input modes available and adapt by transitioning between them with minimum interaction from the user. Users should have the option to override or to require confirmation for mode changes. | 3
| 8.2.6 | The system should accommodate different modes of data input e.g. allow a number of different vocabulary approaches (language systems) including alphabetic, schematic and visual scene (assisted reality – eg bottle of water on table – system ids water) | 1
| 8.2.8 | The system should be able to accommodate users with excessive movement and difficulty moving. | 2
| 8.2.9 | The system should be adaptable for color-blind users or vision issues (e.g. lazy eye) | 3
| 8.2.14 | The system should allow a setting that allows users to use the user interface even if it is difficult to read (e.g. because it has been minimized for maximum see-through capability or because of low sight conditions - an example might be that the UI auto magnifies buttons as they are selected/”moused over”, such as creating a larger dwell target, when the cursor rests on the “a” button, the button enlarges and “pops out” by a set percentage, as selected in a setting menu) | 2
| 8.3.3 | The system should employ assistive aids for target selection (e.g. how “sticky” the cursor is as it moves from target to target). | 1
| 8.3.10 | The system should allow the user to enable screen reading capabilities of navigation elements to assist with low vision users. (e.g. if cursor dwells over “settings”, the system announces “settings”), ideally different from the text-to-speech voice. | 3
| 8.10.2 | The system should alert the user that imput mode is changing based on user’s fatigue with current mode. | 3
| 8.3.7 | The system should adapt to user state by transitioning between input modes (e.g. based on historical user capabilities with head pose, if the user’s selection mode becomes slower, offer a switch to eye tracking). Users should have the option to override or to require confirmation for mode changes.If the system transitions between input modes, it should allow the user to confirm or over-ride the transition. Whether confirmation is required should be adjustable in settings | 3
| URD-U4 | The user should be able to set up the headset for their maximum efficiency & effectiveness
URD-U07 |  | The user should be able to personalize how they interact with the user interface of the device
| 8.2.7 | The system should allow the user to customize the layout of the internal screen | 2
| 8.2.12 | The system should provide an opportunity for the user to rest between selections (e.g. empty space) and pause providing input altogether. | 1
| 8.3.4 | The system should allow the user to control display elements such as font size, speed of cursor, selection method and assistive aids. | 1
| 8.3.11 | The system should have simplified settings menu, easily configured by end users or caregivers. | 2
| 8.4.10 | The system should provide a variety of keyboard layouts (e.g. ABC, QWERTY, Linotype, etc.) | 2
| 8.10.8 | The system should allow user to see and change as many standard system settings as possible (e.g. volume level, system brightness, font size, cursor and dwell color, dwell type, etc.) | 1
URD-U08 |  | The user should be able to compose utterances at least as efficiently as their current system
| 8.2.13 | The system should leverage the user’s entire visual field capacity, preferably adjustable by the user. | 2
| 8.2.15 | The system should follow best practices for 3D UI design (e.g. show examples of the setting change while in the setting menu such as font changes as font is adjusted without requiring the user to exit out of the settings menu) | 2
| 8.3.1 | The system should help the user understand where they are “pointing” and what is being selected | 1
| 8.3.2 | The system should be able to provide visual, auditory and/or haptic feedback to indicate key information (e.g. such as knowing that a selection is being made and/or has been selected.) The type of feedback provided should be adjustable in settings. | 1
| 8.3.5 | The menu and display elements should consistently be in front of the user’s eyes. | 1
| 8.3.8 | The system should achive desired outcomes with a minimum of user interaction (effort). (e.g. minimize the number of selections– preferably under 3) | 1
| 8.3.9 | The system should allow the user to quickly and easily activate re-centering or re-calibration. | 1
| 8.3.12 | System layout and configuration should be accessible prior to acquisition of the physical device. | 2
| 8.4.11 | The system should generate speech with a minimum of “interactions” (e.g. space between words should be added automatically when a word is selected or the system is aware that a word has been created) | 2
| 8.5.3 | The system should provide predictive words and phrases that adjust with each additional input (of a letter or word). | 2
| 8.10.3 | The system should alert the user if re-centering or re-calibration is needed. | 2
| 8.10.7 | The system should indicate battery and internet connectivity status | 1
| URD-U5 | The user should be able to use the headset to maximise their personal agency
URD-U09 |  | User should be able to control how utterances are created, edited, organized
| 8.5.10 | The system should be able to accept and add new “words” to the dictionary (e.g. ability to add unusual names and have them be predicted). | 1
| 8.5.12 | The system should enable an auto-correct feature, to be turned on or off by the user. | 2
URD-U10 |  | User should be able to control how utterances are shared and vocalized
| 8.5.6 | The system should allow the user to receive auditory feedback in either “private” mode for their own purposes (e.g. vocalizing typing letter by letter or commands to alexa), or “public” mode for communicating with others. | 3
| 8.4.2 | The user should be able to specify whether speech is generated as each letter, word or phrase is created. | 1
| 8.4.18 | The system should allow the user to enter a “speaker mode” wherein they can compose and save long segments of speech that could be played back in the form of a speech, and paused periodically to enter “conversation mode”. | 3
| 8.4.19 | The system should allow the user to enter a “document mode” wherein they can compose and save long segments of speech that could be saved and sent via electronic means in the form of texts, journals or letters. | 3
| 8.6.5 | The system should enable the user to send messages electronically – at a minimum via text. | 3
URD-U11 |  | The user should be able to control their environment through connected devices (IoT)
| 8.6.6 | The system should be able to provide user connectivity to the internet and entertainment systems (e.g. Netflix) | 3
| 8.12.1 | The system should allow the addition of other applications on the device. | 3
| 8.6.3 | The user should be able to control their environment with the system through internet connected devices (such as alexa) | 1
| URD-U6 | The headset should enable "live" conversation to the fullest extent possible
URD-U12 |  | Device should be able to adjust predicted words or phrases based on past user inputs
| 8.5.4 | The system’s word prediction should adapt to user utterances & corrections to utterances over time. (e.g. the user increasingly finds the predicted words/phrases relevant & selects them) | 3
| 8.5.7 | The system should allow users to repair/correct sentences and words with minimum “interactions” (e.g. deleting or inserting words in the middle of a sentence or suggesting synonyms/corrections when the cursor hovers over the word). | 1
| 8.5.14 | The system should provide predictive phrases that customarily come before and after each word typed by the user. | 3
| 8.5.1 | The system should provide contextual predictions based on environment (e.g. geographical location, weather, activity). | 3
| 8.5.2 | The system should provide contextual predictions based on utterance history (for example, if a user is talking about a trip to Hawaii, ideally the predictions will tend to enable continued discussion on that topic). | 3
URD-U13 |  | The system should provide quick access or replay of common and context aware words and phrases
| 8.4.4 | The speech generated should be repeatable with a single action (i.e. the phrase is repeated – visually and/or auditorily – every time the play button is pressed) | 1
| 8.4.7 | The system should support pre-recorded utterances for later use. | 2
| 8.4.14 | The system should include a method for rapidly accessing pre-defined user phrases that can ideally be played with a single click without removing the composed message in the main window. (e.g. a phrase might be “please give me a minute to compose my reply”) | 2
| 8.4.15 | The system should allow user to organize their phrase inventory and create shortcuts. | 2
| 8.4.16 | The system should allow users the option of auto clearing speech that has already been generated. | 2
| 8.5.5 | The system should provide a number of “standard phrases”, configurable by the user – including comfort questions, documenting their journey and end of life planning. | 2
URD-U14 |  | The system should allow user to view and interact with others
| 8.1.12 | The system should enable the user to see their surroundings as well as the graphical interface. | 1
| 8.3.6 | The menu and display elements should easily be toggled on and off screen to allow user to look at the outside world unobstructed. | 2
URD-U15 |  | The user should be able to vocalize and display their utterances
| 8.1.7 | The displayed text on the outside of the screen should be legible to people interacting with the user in an indoor and outdoor setting. | 2
| 8.1.8 | The audio should be adjustable to be loud enough to be heard in an indoor and outdoor setting. | 2
| 8.4.3 | The speech generated should be loud enough to be heard in an indoor and outdoor setting(e.g. in a restaurant. Ideally consistent with human dynamic range.) | 1
| 8.4.12 | The system should allow speech to be generated visually and/or auditorily – as preferred by the user. | 1
URD-U16 |  | The system should be able to provide indication that the user is composing utterances
| 8.4.13 | The system should be able to indicate both visually and auditorily, per user preference, that the user is composing speech. (e.g. Please be patient while I compose my message.) | 1
| 8.4.17 | The system should enable to pause vocalization during readback and edit to react to interruptions or discussion. (The intent is to allow users to engage in natural conversations – e.g. turn taking, indicating a “let me finish” or “good point”) | 3
URD-U17 | URD-U7 | The headset should accommodate the user in their own style of speaking and timbre of their utterances
| 8.4.1 | The speech generated should be adjustable in volume and speed. | 1
| 8.4.5 | The system should allow for user personalization of synthesized voice (e.g. banking of voices) before a system is assigned or used. | 3
| 8.4.6 | The system should enable the user to “bank” their voice in an easy or enjoyable manner. (e.g. passive capturing of the voice or capturing the voice via a game) | 3
| 8.4.8 | The system should provide a variety of selectable standard differentiated voices (e.g. differentiated pitches, speeds and resonance) that can be selected as default and used regardless of connectivity status of the system. | 2
| 8.4.9 | The system’s synthesized voice should enable personalization with the user’s desired voice. | 3
| 8.5.11 | The system should generate speech composed within 0.5 seconds of the user pressing “play” and allow the user to increase that delay in settings. | 2
| 8.5.13 | The system should be able to translate composed phrases into different languages using internet translation systems. | 3
| 8.6.4 | The system should be able to leverage existing voice banking repositories (e.g. SAPI5) | 2
| URD-U8 | The User should be able to create and store longer compositions for utterance or digital transmission (e.g. in social media)
URD-U18 |  | Users should not be without a headset for more than two weeks due to service needs
| 8.8.1 | The system should be serviceable by the company. | 2
| 8.8.2 | If service is needed, the user should not be left without a speech generating system for longer than one (1) week. | 2
| 8.8.3 | The system should provide a 1 year manufacturer’s warranty and support. | 2
| 8.8.4 | The company should provide extended warranty and support options. | 3
| 8.8.5 | The system should include a number of field replacable parts for those items of the design most likely to break, such that it maximises the time that the user can use the system. | 1
| URD-C1 | The Caregiver should easily be able to troubleshoot and streamline the user's device
URD-C01 |  | The Caregiver should be able to troubleshoot the user's device without taking the device from the user
| 8.7.1 | The system should provide a method to send a message to the company requesting customer support. | 2
| 8.7.2 | The system should provide a method to receive real-time customer support from the company. | 3
URD-C02 |  | The Caregiver should be able to customize the user’s interface without taking the device from the user
| 8.7.6 | The system should provide speech generation performance data for the user’s circle of care (e.g. software/access settings to make the speech generation easier, faster and better or more efficient for the user, ie. change the settings on the dwell time if they are becoming tired). | 3
| 8.7.10 | The system should allow clinicians and caregivers to change and organize saved phrases in an efficient manner. | 2
| URD-C2 | The Caregiver should be able to confirm that the device is working as expected without removing the device from the User
URD-C03 |  | The Caregiver should be able to confirm that the device is set up and operating correctly without removing the device from the User
| 8.7.4 | The system should enable the user’s circle of care to see what the user is seeing from another device. | 2
| 8.7.5 | The system should enable the user’s circle of care to control the the user’s system from another device. | 3
| 8.7.9 | The system should provide an “admin” mode for clinicians and caregivers to customize the experience for the user. | 3
| URD-C3 | The Caregiver should be able to correctly apply the device with a minimum of adjustments
URD-C04 |  | The Caregiver should be able to correctly apply the device to the user
| 8.7.11 | The system should include “how-to” videos to assist with basic operation of the unit. | 2
| 8.7.12 | The company should provide training and tools to enabe the user, caregivers and clinicians to operate the device. | 1
| URD-M1 | Cognixion should be able to remotely support users and caregivers in using and troubleshooting their device
URD-M01 |  | Cognixion should be able to remotely support users and caregivers in using and troubleshooting their device
| 8.7.3 | The system should allow the company to provide technical support in a virtual manner (e.g. ability to “take over” the headset UI to trouble shoot/repair/train). | 3
| 8.7.7 | The system should provide more-than-the-minimum required training of the user within the device. | 2
| 8.7.8 | The system should provide training of the user in a way that is enticing for the user (e.g. via games vs wall-of-text) | 3
| 8.8.6 | The system should include “how-to” videos to assist with basic maintenance and simple repairs of the unit. | 2
| 8.9.3 | The system should send feedback on errors and bugs to the company while protecting the user’s identity. | 3
| URD-M2 | Cognixion should be able to meet the requirments of standard regulatory agencies such as CMS and the FDA
URD-M02 |  | The device should meet or exceed key requirements of CMS and the FDA
| 8.5.8 | The system shall be usable by Spanish speakers in the United States. | 3
| 8.5.9 | The system shall be usable by English speakers in the United States. | 1
| 8.10.4 | The system should provide a means for the user to easily call for help | 1
| 8.10.5 | The system should provide a means for the user to call 911 | 3
| 8.10.6 | The system should alert the user and shut down if it detects excessive heat or other system parameters that could cause harm to the user. | 1
| 8.11.1 | The system shall meet standard regulatory and ISO standards. | 1
| 8.11.2 | The system shall meet standarAAC (Augmentative and Alternative Communication) requirements | 1
| 8.11.3 | The system shall meet requirements for funding | 2
| 8.11.4 | The system shall meet applicable HIPPA and cybersecurity requirements | 2
| URD-M3 | Cognixion should have a means for addressing customer service requests with minimum cost and maximum efficiency
URD-M03 |  | Cognixion should have means to receive feedback on the system’s performance while the system is in the field
| 8.9.1 | The user and caregiver should be able to provide feedback on the system within the system. | 2
| 8.9.2 | The system should enable the tracking of use metrics that protect the users identiy. | 3

## MCRD (hidden)
URD # | Req # | Requirement Description | Priority (1,2,3) | PRD #
8.1
| 8.1.1 | The system should be comfortable to wear for 1 – 2 hours. | 1
| 8.1.2 | The system should be comfortable to wear for 6-8 hours. | 3
| 8.1.3 | The system should fit at least 95% of average adult heads (i.e. within normal limits). | 2
| 8.1.4 | The system should have a self-contained battery life of 2 hours | 1
| 8.1.5 | The system should provide an extended battery option for 8 hours of standard use. | 3
| 8.1.6 | The screen on the system should be legible to the user in an indoor and outdoor setting. | 2
| 8.1.7 | The displayed text on the outside of the screen should be legible to people interacting with the user in an indoor and outdoor setting. | 2
| 8.1.8 | The audio should be adjustable to be loud enough to be heard in an indoor and outdoor setting. | 2
| 8.1.9 | The system should be comfortable in a seated and reclined position. (e.g. when a patient with ALS wakes up, they should be able to communicate an awakened state to their caregiver) | 2
| 8.1.10 | The system should be comfortable for use in a wheelchair with a headrest. | 1
| 8.1.11 | The use of the headset should be a pleasant and relaxing experience both from a feel of the headset and the design of the user interface | 2
| 8.1.12 | The system should enable the user to see their surroundings as well as the graphical interface. | 1
8.2
| 8.2.1 | The system should allow for input modes of head movement and visual evoked potentials. | 1
| 8.2.2 | The system should allow for input modes of eye tracking | 2
| 8.2.3 | The system should allow the user physical methods such as a joystick or mouse to focus on a visual target (instead of head pose, for example) | 3
| 8.2.4 | The system should enable the use of a switch for quicker selection regardless of input mode being used. | 2
| 8.2.5 | The system should be able to measure user’s ability to use the input modes available and adapt by transitioning between them with minimum interaction from the user. Users should have the option to override or to require confirmation for mode changes. | 3
| 8.2.6 | The system should allow a number of different vocabulary approaches (language systems) including alphabetic, schematic and visual scene (assisted reality – eg bottle of water on table – system ids water) | 1
| 8.2.7 | The system should allow the user to customize the layout of the internal screen | 2
| 8.2.8 | The system should be able to accommodate users with excessive movement and difficulty moving. | 2
| 8.2.9 | The system should be adaptable for color-blind users or vision issues (e.g. lazy eye) | 3
| 8.2.10 | The system should be adaptable for users who wear glasses | 2
| 8.2.11 | The system should be adaptable for users with hearing peripherals | 3
| 8.2.12 | The system should provide an opportunity for the user to rest between selections (e.g. empty space) and pause providing input altogether. | 1
| 8.2.13 | The system should leverage the user’s entire visual field capacity, preferably adjustable by the user. | 2
| 8.2.14 | The system should allow a setting that auto magnifies buttons as they are selected/”moused over”, (e.g. creating a larger dwell target, when the cursor rests on the “a” button, the button enlarges and “pops out” by a set percentage, as selected in a setting menu) | 2
| 8.2.15 | The system should follow best practices for 3D UI design (e.g. show examples of the setting change while in the setting menu such as font changes as font is adjusted without requiring the user to exit out of the settings menu) | 2
8.3
| 8.3.1 | The system should help the user understand where they are “pointing” and what is being selected | 1
| 8.3.2 | The system should be able to provide visual, auditory and/or haptic feedback to indicate key information (e.g. such as knowing that a selection is being made and/or has been selected.) The type of feedback provided should be adjustable in settings. | 1
| 8.3.3 | The system should employ assistive aids for target selection (e.g. how “sticky” the cursor is as it moves from target to target). | 1
| 8.3.4 | The system should allow the user to control display elements such as font size, speed of cursor, selection method and assistive aids. | 1
| 8.3.5 | The menu and display elements should consistently be in front of the user’s eyes. | 1
| 8.3.6 | The menu and display elements should easily be toggled on and off screen to allow user to look at the outside world unobstructed. | 2
| 8.3.7 | The system should adapt to user state by transitioning between input modes (e.g. based on historical user capabilities with head pose, if the user’s selection mode becomes slower, offer a switch to eye tracking). Users should have the option to override or to require confirmation for mode changes.If the system transitions between input modes, it should allow the user to confirm or over-ride the transition. Whether confirmation is required should be adjustable in settings | 3
| 8.3.8 | The system should achive desired outcomes with a minimum of user interaction (effort). (e.g. minimize the number of selections– preferably under 3) | 1
| 8.3.9 | The system should allow the user to quickly and easily activate re-centering or re-calibration. | 1
| 8.3.10 | The system should allow the user to enable screen reading capabilities of navigation elements to assist with low vision users. (e.g. if cursor dwells over “settings”, the system announces “settings”), ideally different from the text-to-speech voice. | 3
| 8.3.11 | The system should have simplified settings menu, easily configured by end users or caregivers. | 2
| 8.3.12 | System layout and configuration should be accessible prior to acquisition of the physical device. | 2
8.4
| 8.4.1 | The speech generated should be adjustable in volume and speed. | 1
| 8.4.2 | The user should be able to specify whether speech is generated as each letter, word or phrase is created. | 1
| 8.4.3 | The speech generated should be loud enough to be heard in an indoor and outdoor setting(e.g. in a restaurant. Ideally consistent with human dynamic range.) | 1
| 8.4.4 | The speech generated should be repeatable with a single action (i.e. the phrase is repeated – visually and/or auditorily – every time the play button is pressed) | 1
| 8.4.5 | The system should allow for user personalization of synthesized voice (e.g. banking of voices) before a system is assigned or used. | 3
| 8.4.6 | The system should enable the user to “bank” their voice in an easy or enjoyable manner. (e.g. passive capturing of the voice or capturing the voice via a game) | 3
| 8.4.7 | The system should support pre-recorded utterances for later use. | 2
| 8.4.8 | The system should provide a variety of selectable standard differentiated voices (e.g. differentiated pitches, speeds and resonance) that can be selected as default and used regardless of connectivity status of the system. | 2
| 8.4.9 | The system’s synthesized voice should enable personalization with the user’s desired voice. | 3
| 8.4.10 | The system should provide a variety of keyboard layouts (e.g. ABC, QWERTY, Linotype, etc.) | 2
| 8.4.11 | The system should generate speech with a minimum of “interactions” (e.g. space between words should be added automatically when a word is selected or the system is aware that a word has been created) | 2
| 8.4.12 | The system should allow speech to be generated visually and/or auditorily – as preferred by the user. | 1
| 8.4.13 | The system should be able to indicate both visually and auditorily, per user preference, that the user is composing speech. (e.g. Please be patient while I compose my message.) | 1
| 8.4.14 | The system should include a method for rapidly accessing pre-defined user phrases that can ideally be played with a single click without removing the composed message in the main window. (e.g. a phrase might be “please give me a minute to compose my reply”) | 2
| 8.4.15 | The system should allow user to organize their phrase inventory and create shortcuts. | 2
| 8.4.16 | The system should allow users the option of auto clearing speech that has already been generated. | 2
| 8.4.17 | The system should enable to pause vocalization during readback and edit to react to interruptions or discussion. (The intent is to allow users to engage in natural conversations – e.g. turn taking, indicating a “let me finish” or “good point”) | 3
| 8.4.18 | The system should allow the user to enter a “speaker mode” wherein they can compose and save long segments of speech that could be played back in the form of a speech, and paused periodically to enter “conversation mode”. | 3
| 8.4.19 | The system should allow the user to enter a “document mode” wherein they can compose and save long segments of speech that could be saved and sent via electronic means in the form of texts, journals or letters. | 3
8.5
| 8.5.1 | The system should provide contextual predictions based on environment (e.g. geographical location, weather, activity). | 3
| 8.5.2 | The system should provide contextual predictions based on utterance history (for example, if a user is talking about a trip to Hawaii, ideally the predictions will tend to enable continued discussion on that topic). | 3
| 8.5.3 | The system should provide predictive words and phrases that adjust with each additional input (of a letter or word). | 2
| 8.5.4 | The system’s word prediction should adapt to user utterances & corrections to utterances over time. (e.g. the user increasingly finds the predicted words/phrases relevant & selects them) | 3
| 8.5.5 | The system should provide a number of “standard phrases”, configurable by the user – including comfort questions, documenting their journey and end of life planning. | 2
| 8.5.6 | The system should allow the user to receive auditory feedback in either “private” mode for their own purposes (e.g. vocalizing typing letter by letter or commands to alexa), or “public” mode for communicating with others. | 3
| 8.5.7 | The system should allow users to repair/correct sentences and words with minimum “interactions” (e.g. deleting or inserting words in the middle of a sentence or suggesting synonyms/corrections when the cursor hovers over the word). | 1
| 8.5.8 | The system shall be usable by Spanish speakers in the United States. | 3
| 8.5.9 | The system shall be usable by English speakers in the United States. | 1
| 8.5.10 | The system should be able to accept and add new “words” to the dictionary (e.g. ability to add unusual names and have them be predicted). | 1
| 8.5.11 | The system should generate speech composed within 0.5 seconds of the user pressing “play” and allow the user to increase that delay in settings. | 2
| 8.5.12 | The system should enable an auto-correct feature, to be turned on or off by the user. | 2
| 8.5.13 | The system should be able to translate composed phrases into different languages using internet translation systems. | 3
| 8.5.14 | The system should provide predictive phrases that customarily come before and after each word typed by the user. | 3
8.6
| 8.6.1 | The system should generate speech even when not connected to the internet. | 1
| 8.6.2 | The system should be able to access banked voices even when not connected to the internet. | 2
| 8.6.3 | The user should be able to control their environment with the system through internet connected devices (such as alexa) | 1
| 8.6.4 | The system should be able to leverage existing voice banking repositories (e.g. SAPI5) | 2
| 8.6.5 | The system should enable the user to send messages electronically – at a minimum via text. | 3
| 8.6.6 | The system should be able to provide user connectivity to the internet and entertainment systems (e.g. Netflix) | 3
| 8.6.7 | They system should enable use of a cell-phone as a hot-spot. | 3
8.7
| 8.7.1 | The system should provide a method to send a message to the company requesting customer support. | 2
| 8.7.2 | The system should provide a method to receive real-time customer support from the company. | 3
| 8.7.3 | The system should allow the company to provide technical support in a virtual manner (e.g. ability to “take over” the headset UI to trouble shoot/repair/train). | 3
| 8.7.4 | The system should enable the user’s circle of care to see what the user is seeing from another device. | 2
| 8.7.5 | The system should enable the user’s circle of care to control the the user’s system from another device. | 3
| 8.7.6 | The system should provide speech generation performance data for the user’s circle of care (e.g. software/access settings to make the speech generation easier, faster and better or more efficient for the user, ie. change the settings on the dwell time if they are becoming tired). | 3
| 8.7.7 | The system should provide more-than-the-minimum required training of the user within the device. | 2
| 8.7.8 | The system should provide training of the user in a way that is enticing for the user (e.g. via games vs wall-of-text) | 3
| 8.7.9 | The system should provide an “admin” mode for clinicians and caregivers to customize the experience for the user. | 3
| 8.7.10 | The system should allow clinicians and caregivers to change and organize saved phrases in an efficient manner. | 2
| 8.7.11 | The system should include “how-to” videos to assist with basic operation of the unit. | 2
| 8.7.12 | The company should provide training and tools to enabe the user, caregivers and clinicians to operate the device. | 1
8.8
| 8.8.1 | The system should be serviceable by the company. | 2
| 8.8.2 | If service is needed, the user should not be left without a speech generating system for longer than one (1) week. | 2
| 8.8.3 | The system should provide a 1 year manufacturer’s warranty and support. | 2
| 8.8.4 | The company should provide extended warranty and support options. | 3
| 8.8.5 | The system should include a number of field replacable parts for those items of the design most likely to break, such that it maximises the time that the user can use the system. | 1
| 8.8.6 | The system should include “how-to” videos to assist with basic maintenance and simple repairs of the unit. | 2
8.9
| 8.9.1 | The user and caregiver should be able to provide feedback on the system within the system. | 2
| 8.9.2 | The system should enable the tracking of use metrics that protect the users identiy. | 3
| 8.9.3 | The system should send feedback on errors and bugs to the company while protecting the user’s identity. | 3
8.1
| 8.10.1 | The system should indicate low battery status when 30 minutes of standard use remain and critical battery status when 10 minues remain. | 2
| 8.10.2 | The system should alert the user that imput mode is changing based on user’s fatigue with current mode. | 3
| 8.10.3 | The system should alert the user if re-centering or re-calibration is needed. | 2
| 8.10.4 | The system should provide a means for the user to easily call for help | 1
| 8.10.5 | The system should provide a means for the user to call 911 | 3
| 8.10.6 | The system should alert the user and shut down if it detects excessive heat or other system parameters that could cause harm to the user. | 1
| 8.10.7 | The system should indicate battery and internet connectivity status | 1
| 8.10.8 | The system should allow user to see and change as many standard system settings as possible (e.g. volume level, system brightness, font size, cursor and dwell color, dwell type, etc.) | 1
8.11
| 8.11.1 | The system shall meet standard regulatory and ISO standards. | 1
| 8.11.2 | The system shall meet standarAAC (Augmentative and Alternative Communication) requirements | 1
| 8.11.3 | The system shall meet requirements for funding | 2
| 8.11.4 | The system shall meet applicable HIPPA and cybersecurity requirements | 2
8.11
| 8.12.1 | The system should allow the addition of other applications on the device. | 3